{"text_format":"3","workbook":"xls","n_sheets":1,"sheets":["公开发布"],"not_rendered":{"external_links":1,"defined_names":2},"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"公开发布","index":0,"n_rows":4,"n_cols":6,"rows_kept":4,"merged":[[0,0,0,5],[1,0,1,5]],"flags":["formulas_cached_only"]}
["附件1"]
["项 目 指 南"]
["序号","项目\n名称","研究目标","技术指标","成果形式","备注"]
[1,"双聚焦多接收电感耦合等离子体质谱仪","针对核工业等领域对同位素比值的高精度分析需求，开展双聚焦多接收电感耦合等离子体质谱仪基础技术攻关、工程化研制和市场化应用，研制具备高性能碰撞反应池的双聚焦多接收电感耦合等离子体质谱仪，完成产品的市场化应用，带动我国该类核仪器达到国际先进水平，解决同位素的准确分析、同质异位素分析等技术问题，提高同位素分析水平，解决该类仪器卡脖子问题。到2030年完成不少于1台双聚焦多接收电感耦合等离子体质谱仪研制，成熟度不低于8级，产品综合成熟度接近100%，产品和技术成熟，经过应用验证，已经实现市场销售，可直接采购。","1.功能指标\n（1）具备同位素比值高精度分析功能；\n2.性能指标\n（1）分析质量范围3～300amu；\n（2）质量分辨率：低分辨>450（对应10%峰谷）、中分辨>5000（对应5%-95%上升沿）、高分辨>10000（对应5%-95%上升沿）；\n（3）灵敏度>40V/ppm 铀（低分辨，1011Ω高阻），>10V/ppm 锂（低分辨，1011Ω高阻），中分辨的灵敏度不小于低分辨的10%，高分辨的灵敏度不小于低分辨的2%；\n（4）系统稳定性优于 50ppm/30分钟；\n（5）长期稳定性：3%（1h，RSD）；\n（6）短期稳定性：1.5%（10min，RSD）；\n（7）天然铀的同位素比（235U/238U）测量内精度优于0.02%，外精度优于0.03%，天然锂的同位素比测量内精度优于0.05%，外精度优于0.06%；\n（8）法拉第放大器噪声小于20µV（1011Ω高阻，积分时间4s）；\n（9）法拉第杯不少于9个，二次电子倍增器不少于2个；\n（10）碰撞反应模式下典型材料的分离效率>99.5%；\n（11）手套箱密封等级优于3级。\n（三）工程化指标。\n（1）成熟度不低于8级，产品综合成熟度接近100%，产品和技术成熟，经过应用验证，已经实现市场销售，可直接采购。","1.实现不少于1台/套（含应用软件）的产品销售（销售合同）；\n2.标准、规范不少于2项；\n3.申请并受理发明专利不少于2项。"]